Instruction    Manual   for    Oscillating

Flow

Version 1.5.0

ibidi instruction manual for oscillating flow

ii

© ibidi GmbH 2011, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

1   Instructions for  oscillating flow experiments

In order  to  apply oscillating  flow to  the  µ-Slides you  need at  least  two Fluidic  Units (one

‘master’ and  one ‘slave’  Fluidic Unit) to  separate the  switching times of  the two  valves of

the Fluidic Unit as simultaneous switching results in unidirectional flow. The basic principle is

such that  the master  Fluidic Unit has  a long  switching time tmaster.  During tmaster  the master

Fluidic Unit  supplies an unchanged  air flow to  the reservoirs of  the slave Fluidic  Unit. Now

the switching time  of the slave Fluidic Unit  tslave can be set as  a fraction of the  tmaster so that

the flow direction  is reversed tmaster /  tslave times before the  master Fluidic Unit  switches the

air  flow to  the  reservoirs. As  a  result you  have  a setup  were  one  Fluidic Unit  creates  a

unidirectional flow and  one which supplies oscillating  flow within the µ-Slide channel.  Since

only one controlling master Fluidic Unit is needed but the pump can control up to four  Fluidic

Units, you can extend the setup to operate with up to 3 oscillating slave Fluidic Units.

1.1   Setting up the Fluidic Units

You   will   also  need   air   tubing   splitters   (Figure   1),   besides   the   two  Fluidic   Units.

Figure 1: Air pressure splitter for 2 Fluidic Units

ibidi instruction manual for oscillating flow

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

1

Pressure IN

Pressure

OUT

Pressured

air

connector

to pump

Valve

cable

connection

Figure 2: Fluidic Unit connectors

Please follow the steps below for correct installation and compare with Figure 2:

1)   Connect the air pressure tube of the pump to the master Fluidic Unit.

2)   Use one of the splitters, for example, the yellow marked to connect the left ‘Pressure

OUT’ port of the Master Fluidic Unit to

a)   the left ‘Pressure IN’ port of the master Fluidic Unit and

b)   the left ‘Pressure IN’ port of the slave Fluidic Unit.

3)   Repeat step 2) with the black marked air splitters and the right side of the master and

slave Fluidic Unit.

4)   Connect  the  pump   and  the  two   Fluidic  Units  with  the   electric  cables.  In   the

presented case we  use ‘Port 1’ for  the master Fluidic Unit  and ‘Port 2’ for  the slave

Unit.

ibidi instruction manual for oscillating flow

2

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

Figure 3: Connected master (MFU) and slave Fluidic Unit (SFU)

The setup should like like Figure 3

1.2   Oscillating experiment with 4 Fluidic Units

To  extend  an   experiment  to  one  master   Unit  and  3  oscillating  slave   Fluidic  Units,  a

corresponding air pressure splitter  is needed. The connection is  according to the setup with

two Fluidic Units. Figure 4 shows an example of the required air pressure tubing.

The connection is  according to section 1.1 with  a simple extension for  the additional Fluidic

Units. The connection to the air pressure tubing is shown in Figure 5.

ibidi instruction manual for oscillating flow

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

3

Figure 4: Air tubing for experiment with 4 Fluidic Units: 1 master Fluidic Unit (MFU) and 3 slave

Fluidic Units (SFU 1-3).

ibidi instruction manual for oscillating flow

4

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

Figure 5: Air pressure connection with a setup using 4 Fluidic Units

1.3   Operation mode for oscillation flow experiments

Possible?

Oscillating flow

Continuous flow

Master Fluidic Unit   (FU 1)

no

yes

yes

yes

yes

Slave Fluidic Unit

Slave Fluidic Unit

Slave Fluidic Unit

(FU 2)

(FU 3)

(FU 4)

yes

yes

yes

Please  be  aware  that  the  master  Fluidic  Unit  can  only  be  used  for  unidirectional  flow

experiments. All slave  Fluidic Units can  reverse the flow direction  with switching times  set

by PumpControl.

1.4   Settings within PumpControl

As  the   switching  times  are  different   for  the  master   and  the  slave   Fluidic  Units  the

PumpControl program  of the ibidi  Pump has  to be set  accordingly. For that  you find in  the

PumpControl  software checkboxes  for  ‘continuous’  and ‘oscillating’  valves  and switching

operations. Please check Figure 6 how to correctly set the corresponding parameters.

ibidi instruction manual for oscillating flow

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

5

Pump port for

continuous flow:

master Fluidic

Unit

Slow switching time

for the master Fluidic

Unit: tmaster

Fast switching time

for the slave Fluidic

Unit: tslave

Pump ports for

oscillating flow:

Slave Fluidic

Units

Figure 6: Screen shot of settings for oscilalting experiments

1.5   Equilibrating master and slave Fluidic Units

As master and slave Fluidic  Unit are connected to the same air  pressure and switching time

equilibration of the liquid levels of  both Fluidic Units cannot be done simultaneously. To stop

the  liquid movement  in  the  reservoirs of  either  Fluidic Unit  the  Perfusion  Set has  to  be

clamped off  (pinch off  test described in  instruction manual  for the pump  system) with  the

supplied  hose clamp.  As  the order  is  indifferent  you can  start  off equilibrating  the  liquid

levels  of  the  master  Fluidic  Unit  by  clamping  off  the  Perfusion   Set  of  the  slave  unit.

Afterwards the same should  be done with the slave  unit. This step is beneficial  as it shows

if the Perfusion Set is mounted correctly in both Fluidic Units.

ibidi instruction manual for oscillating flow

6

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

ibidi instruction manual for oscillating flow

© ibidi GmbH 2012, version 1.5.0, revision 1, 2012-08-03, Dr. Miro Kroutvar

7