Instructions

µ–Slide2Well Ph+GlassBottom

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhave

allbeendesignedforhigh–endmicroscopicanalysisofﬁxedorlivingcells.

Theglassbottomversionsoftheµ–Slidesandµ–DishesareespeciallydesignedforTIRF

andsinglemoleculeapplications.

Theµ–Slide2Well   Ph+GlassBottom(Phasecontrastplus)isanarrayof2squareﬁelds

wherecellscanbecultivatedandinvestigatedwithmicroscopicalmethods.     Theµ–Slide

2WellPh+GlassBottomimprovestheopticalqualityofphasecontrastmicroscopy.Incon-

trasttotheclassicµ–Slide2Well,thePh+versionprovidesaspecialplateinthecenterof

thewells.Thisplatesuppressesthemeniscuswhichisdisturbingthephasecontrasteffectinnormalopenwells.Openings

nearthecornersprovideaccesstothewellsforﬁllingandaspiratingliquidseasily.

Material

ShelfLifeofDifferentSurfaces

ibiTreat,GlassBottom,ESS

Collagen,Poly-Lysine

Fibronectin

36months

18months

4months

The glass bottom version of the µ–Slides are made of a

standardµ–Slidebutwithaglasscoverslipbottom.     Itis

not possible to detach the bottom. The µ–Slides are not

autoclavablesincetheyaretemperaturestableonlyupto

80°C/175°F.

Geometry

OpticalPropertiesibidiGlassBottom

Theµ–Slide2Well   Ph+GlassBottomprovidesastandard

slideformataccordingtoISO8037/1.

RefractiveindexnD

Abbenumber

Thickness

1.523

55

No.1.5H(selectedquality

170µm,±5µm)

Geometryofµ–Slide2WellPh+GlassBottom

Numberofwells

2

Material

Schottborosilicateglass,

D263M

Dimensionsofwells(w×l×h)inmm     21.2 × 23.3 ×

3.0

Growthareaperwell

Coatingareaperwell

Volumeperwell

4.8cm2

11.4cm2

Attention!

1.5ml

Be cautious when handling µ–Slides and µ–Dishes

with glass bottom! The glass coverslip is very frag-

ileandmightbreakeasily.  Handlewithcaretoavoid

physicalinjuryanddamagetodevicesthroughleak-

ageofthemedium.

Liquidheight

3.0mm

10.8mm

No.1.5

Totalheightwithlid

Bottommatchescoverslip

Surfaceandcoating

Theµ–Slide2WellPh+GlassBottomismanufacturedwith

ShippingandStorage

anuncoatedglasscoverslip.Washingsteps(e.g.withPBS)

beforecellseedingcanremoveglassdustwhichisadvan-

tageousfordirectcellgrowthonthesurface.

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

Protein coatings increase direct cell growth of adherent

cells.Speciﬁccoatingsonglassarepossiblefollowingthis

protocol:

Conditions

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.  Prepareyour

µ–Slide.Adjusttheconcentrationtoacoatingareaof

11.4cm2and1.5ml.

Shippingconditions

Storageconditions

Ambient

RT(15-25°C)

µ–Slide2WellPh+GlassBottom

Page1

Version2.0(2015-11-22)

Instructions

µ–Slide2Well Ph+GlassBottom

• Apply1.5mlintothegrowtharea.    Makesurethat

theentirebottomiscoveredwithliquideasilytilting

orshakingtheµ–Slide.   Putonthelidandleaveat

roomtemperatureforatleast30minutes.

FillingandHandling

• Aspirate the solution and wash with the recom-

mendedproteindilutionbuffer.   Optionally, letdry

atroomtemperature.  Attention, somecoatingpro-

teinsmightdegeneratewhendrying!

Fill the wells by using a standard pipet. Inject the cell

suspensiondirectlyintooneoftheopenings.Mediumex-

changeiseasilydonebyaspiratingtheentirevolumeand

reﬁllingusing1.5mlperwell.

SolventsforFixationandStaining

Cellscanbeobservedliveorﬁxeddirectlyintheµ–Slide

preferablyonaninvertedmicroscope.Theslidematerialis

compatibletoacids,alkalis,PFA,andsiliconeoil.Alcohols

maybeusedforshorttermincubation(e.g.    cellﬁxation).

Acetoneisnotcompatible.   Furtherspeciﬁcationscanbe

foundat

For optimal results in ﬂuorescence microscopy and stor-

ageofstainedprobesibidiprovidesamountingmedium

(50001)optimizedforµ–Dishesandµ–Slides.

SeedingCells

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,applicationofa5-11×104cells/ml

suspensionshouldresultinaconﬂuentlayerwithin

2–3days.

• Apply1.5mlcellsuspensionintoeachwelloftheµ–

Slide.Avoidshakingasthiswillresultininhomoge-

neousdistributionofthecells.

ImmersionOil

• Cover the slide with the supplied lid. Incubate at

37°Cand5%CO2asusual.

Undemanding cells can be left in their seeding medium

foruptothreedaysandgrowtoconﬂuencethere.     How-

ever, best results might be achieved when the medium

is changed every 1–2 days. Carefully aspirate the old

medium and replace it by 1.5ml/well fresh medium.

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

Tip:

Thedaybeforeseedingthecellswerecommendplac-

ingthecellmediumandtheµ–Slideintotheincubator

forequilibration. Thiswillpreventtheliquidinside

fromemergingairbubblesovertheincubationtime.

Company

Product

OrderingNumber

Zeiss

Zeiss

Leica

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

µ–Slide2WellPh+GlassBottom

Page2

Version2.0(2015-11-22)

Instructions

µ–Slide2Well Ph+GlassBottom

µ–Slide2well Ph+selectionguide

µ–Slide2Well

µ–Slide2WellPh+

Standard open wells for maxi-  Specialplateinthecenterofthe

mum sample access.

Meniscus

wells suppresses meniscus for-

mation. Nomeniscus–parallel

disturbs the beam path.

Good

phasecontrastqualityonlyinthe    beampath.  Forexcellentphase

centerofeachwell

contrastmicroscopyalloverthe

wells.

µ–Slide2WellPh+GlassBottom

Page3

Version2.0(2015-11-22)

Instructions

µ–Slide2Well Ph+GlassBottom

OrderingInformation

Theµ–Slide2WellisavailableasopenwellandasaPh+version,aswellasinaglassbottomversion.          Seetablebelowfor

choosingyourµ–Slide2Well.

µ–Slide2Well

Cat.No.

Description

80886

80882

80883

80884

80885

80881

80887

µ–Slide2WellibiTreat:#1.5polymercoverslip,tissueculturetreated,sterilized

µ–Slide2WellCollagenIV:#1.5polymercoverslip,sterilized

µ–Slide2WellFibronectin:#1.5polymercoverslip,sterilized∗

µ–Slide2WellPoly-L-Lysine:#1.5polymercoverslip,sterilized

µ–Slide2WellPoly-D-Lysine:#1.5polymercoverslip,sterilized∗

µ–Slide2WellUncoated:#1.5polymercoverslip,hydrophobic,sterilized

µ–Slide2WellGlassBottom:1.5H(170µm±5µm)D263MSchottglass,sterilized

∗availableonrequestonly

µ–Slide2WellPh+

Cat.No.

Description

80296

80292

80293

80294

80295

80291

80297

µ–Slide2WellPh+ibiTreat:#1.5polymercoverslip,tissueculturetreated,sterilized

µ–Slide2WellPh+CollagenIV:#1.5polymercoverslip,sterilized

µ–Slide2WellPh+Fibronectin:#1.5polymercoverslip,sterilized∗

µ–Slide2WellPh+Poly-L-Lysine:#1.5polymercoverslip,sterilized

µ–Slide2WellPh+Poly-D-Lysine:#1.5polymercoverslip,sterilized∗

µ–Slide2WellPh+Uncoated:#1.5polymercoverslip,hydrophobic,sterilized

µ–Slide2WellPh+GlassBottom:1.5H(170µm±5µm)D263MSchottglass,sterilized

∗availableonrequestonly

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–Slide2WellPh+GlassBottom

Page4

Version2.0(2015-11-22)