Instructions

micro–Insert4Well

Themicro–Insert4Wellisaproductforlongtermmicroscopystudiesespeciallystemcellcultivation.

Aspecialstickyandbiocompatiblesurfaceatthebottomsideworkslikeaglueandavoidsleaking.

Adherent or non–adherent cells placed into the small wells sink to the bottom and can be analyzed

microscopically. Thewellsproviderectangulardigitalformat(4:3)forcoveringtheentireareabyCCD

cameras on microscopes. Inclined walls provide excellent optical access to all cells located near the

edgesofthewell.Optionally,aftercellattachmenttheinsertcanberemovedbyusingsteriletweezers.

Therearenoremainsonthesurface.

Material

SeedingCells

• Preparecellsuspensionasusual.Dependingonyour

Themicro–Insert4Wellismanufacturedfrombiocompat-

ible silicone material. Although, the material is auto-

clavableandcompatibletoalcoholswedonotrecommend

reusingit.

celltypeapplicationofa1-7×105cells/mlshould

resultinaconﬂuentlayerwithin24hours.

• Apply 10 µl into each well. Avoid shaking as this

willresultininhomogeneouscelldistribution.

Pleasenote! Whenusingaibidiµ–Dish,µ–Slideorµ–

Platetakecarethattheibidistandardbottomiscompat-

iblewithcertaintypesofimmersionoilonly.     Alistof

suitableoilscanbefoundonpage2.

• Incubateat37°Cand5%CO2asusual.

• Optionally,itispossibletoﬁlltheouterareawithcell

suspensionorcellmedium.

• Conductyourexperiment.

Geometry

Geometryofthemicro–Insert4Well

Numberofwells

Welldimensions

Growthareaperwell

Coatingareaperwell

Diameter

4

2.0mm×1.5mm

0.03cm2

0.23cm2

12mm

4.2mm

10µl

Height

Volumesmallwell

Volumecompleteinsert

150µl

Werecommendusingthemicro–Insert4Wellinibidiµ–

Dishes,µ–Slide2Well,6wellplates,12wellplatesorlarge

Petridishes. Itisalsopossibletousethemonsterileglass

coverslipsorglassslides.

Tip:

In case the cell lawn is (partially) removed together

withtheinsertuseasmallerseedingdensitytocreate

alessconﬂuentcelllayerordecreaseincubationtime.

SurfacesandCoatings

Theinsertscanbetransferredtoanyﬂat,     clean, anddry

surface. Usesteriletweezersfortransferandgentlypush.

Keepinmindthatonlythebottomsideissticky.      micro–

Insert4Wellisnotworkingonwetormoistsurfaces.      It

mightalsonotworkonunevenordustysubstrates.

micro–Insert4Well

Page1

Version2013-01-10

Instructions

micro–Insert4Well

ShippingandStorage

ImmersionOil

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

Conditions

Company

Product

OrderingNumber

Shippingconditions

Storageconditions

Ambient

ibidi

Zeiss

Zeiss

Leica

ImmersionOil

(ibidi)50101

RT(15-25°C)

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

ShelfLifeofDifferentSurfaces

ibiTreat,glassbottom,ESS

36months

18months

4months

Collagen,Poly-Lysine

Fibronectin

micro–Insert4WellFamily

Themicro–Insert4Wellisavailableindifferentproductversions.

micro–Insert4Wellinµ–Dish35mm,high

Cat.No.

Description

Characteristics

80406

micro–Insert4Wellinµ–Dish35mm,highibiTreat:readytouse      sterilized,hydrophilic

25micro–Inserts4Wellforselfinsertion

Cat.No.

Description

Characteristics

80409

25micro–Inserts4Wellforselfinsertion:ina10cmtransport

dish

sterilized

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

micro–Insert4Well

Page2

Version2013-01-10