Instructions

µ–Slide2×9well

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhaveallbeen

designedforhigh–endmicroscopicanalysisofﬁxedorlivingcells.     Thehighopticalqualityofthe

materialissimilartothatofglass,soyoucanperformallkindsofﬂuorescenceexperimentswith

uncompromisedresolutionandchoiceofwavelength.Theµ–Slide2×9wellharbourstwoarraysof

3×3squareﬁeldswherecellscanbecultivatedandinvestigatedwithmicroscopicalmethods.       Itis

intendedforcheckingoutexperimentalparameterslikeantibodydilution,seedingdensityormost

effectivedrugconcentrations.  Aspecialapplicationistheco-cultivationofdifferentcelltypeswhichshareonemedium

supernatantbutgrowseparatelyfromeachother.

Material

pensabletotreattheuncoatedµ–Slidewithbiopolymers,

whichmediatecelladhesionandgrowth.

ibidiµ–Slides,µ–Dishes,andµ–Platesaremadeofaplas-

tic that has the highest optical quality. The bottom ma-

terial exhibits extremely low birefringence and autoﬂuo-

rescence, similar to that of glass. Also, it is not possible

todetachthebottomfromtheupperpart.     Theµ–Slides,

µ–Dishes, and µ–Plates are not autoclavable, since they

areonlytemperature–stableupto80°C/175°F.Pleasenote

that gas exchange between the medium and incubator’s

atmosphere occurs partially through the plastic bottom,

whichshouldnotbecovered.

Geometry

Theµ–Slide2×9wellprovidesstandardslideformatac-

cordingtoISO8037/1.

Geometryofµ–Slide2×9well

Numberofmajorwells

Majorwells(w×l×h)inmm

Volumepermajorwell

2

21.5×23.6×6.8

600µl

OpticalPropertiesibidiStandardBottom

RefractiveindexnD(589nm)

Abbenumber

1.52

Numberofminorwells

Minorwells(w×l×h)inmm

Volumeperminorwell

2×9

56

6.1×6.8×1.3

Thickness

No.1.5(180µm)

microscopyplastic

70µl

0.4cm2

0.55cm2

Material

Growthareaperminorwell

Coatingareaperminorwell

Totalheightwithlid

Please note! The ibidi standard bottom is compatible

withcertaintypesofimmersionoilonly.    Alistofsuit-

ableoilscanbefoundonpage2.

8mm

Bottommatchescoverslip

No.1.5

Coatingyourµ–Slide2×9well

µ–SlideSurfaces

Theuncoatedµ–Slidemustbecoatedtopromotecellad-

hesion.Ifyouwanttoestablishacertaincoatingtomatch

your needs, we recommend testing your coating proce-

dure on both uncoated and ibiTreat µ–Slides, since we

haveobservedthatsomebiomoleculesadheredifferently

tohydrophobicandhydrophilicplasticsurfaces.

Dependingonthetypeofcellsandthespecialapplication

youareusing,youwillneedµ–Slideswithdifferentsur-

faces.Ifyoudonotrequireanyspecialadhesionmolecules

foryourapplication,thebestchoicewillbeibiTreat,atis-

sueculturetreatedsurface.

Weprovideprecoatedµ–Slideswithadhesionsubstrates

likeCollagenIV,Fibronectin,Poly–L–Lysin,andPoly–D–

Lysin.Suchadhesionsubstrateshavebeenshowntostim-

ulatetheadhesionandgrowthofvariouscelllinesinµ–

Slides.Onlyhigh–qualitysubstratesareused1.

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.

• Apply70µlperminorwell(=1.26mlperSlide)and

leaveatroomtemperatureforatleast30minutes.

• Aspiratethesolutionandwashwithultra-purewa-

ter.Letdryatroomtemperature.

Theuncoatedµ–Slideismanufacturedfromhydrophobic

plastic. For the cultivation of most cell lines, it is indis-

1CollagenIV:Corning#356233,Fibronectin:Corning#354008,Poly–L–Lysin:Sigma#P4832,Poly–D–Lysin:Corning#354210

µ–Slide2×9well

Page1

Version2014-05-14

Instructions

µ–Slide2×9well

Seedingcells

PreparationforCellMicroscopy

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,applicationofa5–10×104cells/ml

suspensionshouldresultinaconﬂuentlayerwithin

2–3days.

To analyze your cells, no special preparations are neces-

sary.Cellscanbeobservedlive,orﬁxeddirectlyintheµ–

Slideonaninvertedmicroscope.  Youcanuseanyﬁxative

ofyourchoice.  Theµ–Slidematerialiscompatiblewitha

varietyofchemicals(e.g.,AcetoneorMethanol).    Further

speciﬁcationscanbefoundatDuetothe

thinbottomofonly180µm,highresolutionmicroscopyis

possible.

• Usethecenterminorwellforrecipientcellsandthe

8outerwellsforfeedercells.

• Apply40-60µlcellsuspensionintoeachminorwell

oftheµ–Slide.  Avoidshakingasthiswillresultin

inhomogeneousdistributionofthecells.

ImmersionOil

• Coverreservoirswiththesuppliedlid.    Incubateat

37°Cand5%CO2asusual.

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

• Aftercellattachementﬁll400-600µlintoeachlarge

reservoir,allowingthecellstosharefactors.

Undemanding cells can be left in their seeding medium

foruptothreedaysandgrowtoconﬂuencythere.     How-

ever, best results might be achieved when the medium

is changed every 1–2 days. Carefully aspirate the old

mediumandreplaceby1.2mlfreshmedium.

Company

Product

OrderingNumber

ibidi

Zeiss

Zeiss

Leica

ImmersionOil

(ibidi)50101

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

Pleasealsoseeour

µ–Slide2×9wellfamily

Theµ–Slide2×9wellfamilyisavailablewithdifferentsurfaces.Seetablebelowforchoosingyourµ–Slide2×9well.

OrderingNumber  TreatmentorCoating

Characteristics

81806

81802

81803

81804

81805

81801

ibiTreat,sterile

hydrophilic,tissueculturetreated

proteincoating

CollagenIV,sterile

Fibronectin,sterile∗

Poly-L-Lysine,sterile

Poly-D-Lysine,sterile∗

proteincoating

biopolymercoating

biopolymercoating

hydrophobic

uncoated,sterile

∗availableonrequestonly

µ–Slide2×9well

Page2

Version2014-05-14

Instructions

µ–Slide2×9well

SelectedReferences

D.G.Buschke,J.M.Squirrell,J.J.Fong,K.W.Eliceiri,andB.M.Ogle.         Celldeath,non-invasivelyassessedbyintrinsic

ﬂuorescenceintensityofNADH,isapredictiveindicatoroffunctionaldifferentiationofembryonicstemcells.Biologyofthe

Cell,2012.doi:10.1111/boc.201100091.

C.Hagen,P.Guttmann,B.Klupp,S.Werner,S.Rehbein,T.C.Mettenleiter,G.Schneider,andK.Gu¨newald.          Correlative

VIS-ﬂuorescenceandsoftX-raycryo-microscopy/tomographyofadherentcells.     JournalofStructuralBiology,  2012.  doi:

10.1016/j.jsb.2011.12.012.

M.Ko¨ttgen,A.Hofherr,W.Li,K.Chu,S.Cook,C.Montell,andT.Watnick.         DrosophilaSpermSwimBackwardsinthe

FemaleReproductiveTractandAreActivatedviaTRPP2IonChannels.PLoSONE,2011.doi:10.1371/journal.pone.0020031.

R.J.Scheubel,J.Holtz,I.Friedrich,J.Borgermann,S.Kahrstedt,A.Navarrete-Santos,R.E.Silber,andA.Simm.         Paracrine

effectsofCD34progenitorcellsonangiogenicendothelialsprouting.    InternationalJournalofCardiology,2008.  doi:10.1016/

j.ijcard.2008.10.009.

µ–Slide2×9well

Page3

Version2014-05-14

Instructions

µ–Slide2×9well

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–Slide2×9well

Page4

Version2014-05-14