Instructions

µ–SlideIII  3in1

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhaveall

beendesignedforhigh–endmicroscopicanalysisofﬁxedorlivingcells.      Thehighoptical

qualityofthematerialissimilartothatofglass,soyoucanperformallkindsofﬂuorescence

experimentswithuncompromisedresolutionandchoiceofwavelength.

Theµ–SlideIII3in1isdesignedforﬂowassayswithdifferentliquidsmergingintoonechan-

nel. Itcanbeconnectedtoapumpandenablesyoutoobservecellsunderswitchableﬂow

conditions. The design allows generating ﬂuid stable concentration proﬁles in the main

channel for e.g. chemotaxis experiments. The microﬂuidic system can generate spatially

andtemporallycontrolledgradientsofchemotacicfactorsbylaminarﬂow.

Material

pensabletotreattheuncoatedµ–Slidewithbiopolymers,

whichmediatecelladhesionandgrowth.

ibidiµ–Slides,µ–Dishes,andµ–Platesaremadeofaplas-

tic that has the highest optical quality. The bottom ma-

terial exhibits extremely low birefringence and autoﬂuo-

rescence, similar to that of glass. Also, it is not possible

todetachthebottomfromtheupperpart.     Theµ–Slides,

µ–Dishes, and µ–Plates are not autoclavable, since they

areonlytemperature–stableupto80°C/175°F.Pleasenote

that gas exchange between the medium and incubator’s

atmosphere occurs partially through the plastic bottom,

whichshouldnotbecovered.

Geometry

Theµ–SlideIII3in1providesstandardslideformataccord-

ingtoISO8037/1.

Geometryoftheµ–SlideIII3in1

Adapters

femaleLuer

60µl

Volumeperreservoir

Numberofchannels

Totalchannelvolume

Heightofallchannels

Widthofchannelsthin/thick

Totalgrowtharea

OpticalPropertiesibidiStandardBottom

3in1

60µl

RefractiveindexnD(589nm)

Abbenumber

1.52

0.4mm

56

1/3mm

1.23cm2

Thickness

No.1.5(180µm)

microscopyplastic

Material

Distanceofscalebars

Bottommatchescoverslip

1mm

No.1.5

Please note! The ibidi standard bottom is compatible

withcertaintypesofimmersionoilonly.    Alistofsuit-

ableoilscanbefoundonpage4.

µ–SlideSurfaces

Dependingonthetypeofcellsandthespecialapplication

youareusing,youwillneedµ–Slideswithdifferentsur-

faces.Ifyoudonotrequireanyspecialadhesionmolecules

foryourapplication,thebestchoicewillbeibiTreat,atis-

sueculturetreatedsurface.

Coatingyourµ–SlideIII   3in1

Weprovideprecoatedµ–Slideswithadhesionsubstrates

likeCollagenIV,Fibronectin,Poly–L–Lysin,andPoly–D–

Lysin.Suchadhesionsubstrateshavebeenshowntostim-

ulatetheadhesionandgrowthofvariouscelllinesinµ–

Slides.Onlyhigh–qualitysubstratesareused1.

Theuncoatedµ–Slidemustbecoatedtopromotecellad-

hesion.Ifyouwanttoestablishacertaincoatingtomatch

your needs, we recommend testing your coating proce-

dure on both uncoated and ibiTreat µ–Slides, since we

haveobservedthatsomebiomoleculesadheredifferently

tohydrophobicandhydrophilicplasticsurfaces.

Theuncoatedµ–Slideismanufacturedfromhydrophobic

plastic. For the cultivation of most cell lines, it is indis-

1CollagenIV:Corning#356233,Fibronectin:Corning#354008,Poly–L–Lysin:Sigma#P4832,Poly–D–Lysin:Corning#354210

µ–SlideIII3in1

Page1

Version2014-05-13

Instructions

µ–SlideIII  3in1

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.

• Apply60µlcellsuspensionintoadapterBoftheµ–

Slide. Quickdispensinghelpstoavoidtrappedair

bubbles.

• Apply60µltoadapterB(handlingseebelow)and

leaveatroomtemperatureforatleast30minutes.

• Coverreservoirswiththesuppliedlid.    Incubateat

37°Cand5%CO2asusual.

• Aspiratethesolutionandwashwithultra–purewa-

ter.Youcanaddthewaterintoonechannelendand

simultaneously aspirate it on the other side. Take

carethatallofthethreechannelsarewashed.

• Await cell attachment in order not to ﬂush out the

cells. Afterwards ﬁll each reservoir with 60 µl cell

freemedium.

• Optionallyletdryatroomtemperature.    Attention,

somecoatingproteinsmightdegeneratewhendry-

ing!

• Connecttheµ–Slidetothepumpandconductyour

perfusionexperiment.

Furtherinformationaboutcoatingsareprovidedin

Important!

Thedaybeforeseedingthecellswerecommendplac-

ingthecellmediumandtheµ–Slideintotheincubator

forequilibration. Thiswillpreventtheliquidinside

thechannelfromemergingairbubblesovertheincu-

bationtime.

FillingandHandling

• AlwaysﬁllthechannelfromadapterBasshownin

thepicture.

• Whenusingapipetmakesureyouplacethetipdi-

rectlyontothesmallchannelsinlet.

ExchangingMedium

• Especiallytheuncoated,hydrophobicchannelcanbe

ﬁlledmucheasierbyusingasmallvolumesyringe

withaLuertip.

Aspirateallfourreservoirsandﬁll60µloffreshmedium

intoreservoirB,whichwillreplacethechannelvolumeby

gravityﬂow. Repeatthisstepthreetimesuntileachreser-

voirisﬁlledwith60µl.

• Makesurethatalladaptersarecompletelyﬁlledbe-

foreLuerconnectorsarepluggedin.    Otherwiseair

bubbleswillbetrapped.

FluidConnectionsandGradientShapes

2in1

SeedingCells

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,applicationofa3–7×105cells/ml

resultsina20%opticalconﬂuentcelllayerafterat-

tachment.

• A seeding density of 1–4 × 106 cells/ml creates a

100% optical conﬂuent cell layer after cell attach-

ment.

Cliff-shapedGradients

Version2014-05-13

µ–SlideIII3in1

Page2

Instructions

µ–SlideIII  3in1

3in1

Concentrationproﬁlescanbevisualizedbyusingaﬂuo-

rescentdyee.g.rhodamine.Switchingtimebetweenstate

1andstate2canrangefromsecondstohours,depending

onthespeedofthecells’response.

The concentration proﬁle  created is always sigmoid

shapedandisdependingonoverallﬂowrateandposition

insidethechannel.

Hill-shapedGradients

Allﬂowratesshouldbeadjustedinawaythatthepoint

ofinﬂectionofthesigmoidisatidenticalpositioninstate

1andstate2.

Cup-shapedGradients

Example Experiment – Cells in Temporally

ControlledGradients

The following example experiment illustrates the idea

howtosetupaswitchablechemicalgradientinthelarge

channelforachemotaxisexperiment.

The longer liquids are in contact next to each other the

smootherthesigmoidshapebecomes.

Sinceﬂowisusedtokeepthegradientsstable,thereisal-

ways a shear stress applied to the cells. Please perform

controlexperimentswiththeexperimentalﬂowratetoex-

cludepolarizationeffectsfromtheﬂowitself.    Flowrates

andcorrespondingshearstresscanbefoundinApplica-

tionNote11onwww.ibidi.com.   Forexample,aﬂowrate

of1ml/minresultsinashearstressof2.3dyn/cm2.

µ–SlideIII3in1

Page3

Version2014-05-13

Instructions

µ–SlideIII  3in1

ImmersionOil

PreparationforCellMicroscopy

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

To analyze your cells, no special preparations are neces-

sary.Cellscanbeobservedlive,orﬁxeddirectlyintheµ–

Slideonaninvertedmicroscope.  Youcanuseanyﬁxative

ofyourchoice.  Theµ–Slidematerialiscompatiblewitha

varietyofchemicals(e.g.,AcetoneorMethanol).    Further

speciﬁcationscanbefoundatDuetothe

thinbottomofonly180µm,highresolutionmicroscopyis

possible.

Company

Product

OrderingNumber

ibidi

Zeiss

Zeiss

Leica

ImmersionOil

(ibidi)50101

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

µ–SlideIII3in1

Page4

Version2014-05-13

Instructions

µ–SlideIII  3in1

SelectedReferences

X.Cheng,M.Joseph,J.A.Covington,T.Dafforn,M.R.Hicks,andA.Rodger.       Continuouschannelﬂowlineardichroism.

RoyalSocietyofChemistry,2012.doi:10.1039/C2AY25513H.

B.Meier,A.Zielinski,C.Weber,D.Arcizet,S.Youssef,T.Franosch,J.O.R¨adler,andD.Heinrich.Chemotacticcelltrapping

incontrolledalternatinggradientﬁelds.PNAS,2011.doi:10.1073/pnas.1014853108.

F.Moseley, J.Halamek, F.Kramer,  A.Poghossian, M.Schoening, andE.Katz.  Enzyme-basedreversiblecnotlogicgate

realizedinaﬂowsystem.RoyalSocietyofChemistry,2014.doi:10.1039/C4AN00133H.

µ–SlideIII3in1

Page5

Version2014-05-13

Instructions

µ–SlideIII  3in1

µ–SlideIII  3in1

Family

Theµ–SlideIII3in1familyisavailablewithdifferentsurfaces.Seetablebelowforchoosingyourµ–SlideIII3in1.

OrderingNumber  TreatmentorCoating

Characteristics

80316

80312

80313

80314

80315

80311

ibiTreat,sterile

hydrophilic,tissueculturetreated

proteincoating

CollagenIV,sterile

Fibronectin,sterile∗

Poly-L-Lysine,sterile

Poly-D-Lysine,sterile∗

proteincoating

biopolymercoating

biopolymercoating

hydrophobic

uncoated,sterile

∗availableonrequestonly

TubeAdapterSet

Fortheconnectionoftheibidiµ–SlidestoanyﬂowsystemsuitableTubeAdapterSetsarealsoavailable.Theyconsistofa

tubing(20cm)withinnerdiameterof1.6mmandadaptersfortheconnectionbetweentheibidiµ–Slide(femaleLuer)and

thetubingofthepumpinuse.

OrderingNumber  ProductName

Characteristics

12pcs,sterile

10831

TubeAdapterSet

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–SlideIII3in1

Page6

Version2014-05-13